## Calculator
ABR Masking Noise Calculator 2024
Instructions:
Click in pale yellow cells to enter options and appropriate levels then press Enter / Return | [12 empty] | <-3
Warning: some variables will be unknown (e.g. air-bone gaps) so enter several likely values; if in doubt use the higher noise level
This calculator is a time-saving guide; the user must carry clinical responsibility for any decisions they make
Result is the dB "dial" level of noise for the selected system, stimulus type, transducer, corrected age & stimulus level | [9 empty] | Label | Freq | 4k | -3 to -1
Offset is an alternative way of specifying noise level: it is the dB noise level relative to the stimulus level | [9 empty] | Click | Wband |  | 0 to 2
[6 empty] | Warning! High Noise Level |  |  |  | 4k pip | 4k |  | 3 to 5
ABR System: |  | Interacoustics Eclipse |  |  |  |  | moderate |  |  | 2k pip | 2k |  | 6 to 8
Stimulus Transducer: |  | BC |  |  |  | Warning! There is a moderate risk of cross-masking |  |  |  | 1k pip | 1k |  | 9 to 12
Noise Transducer: |  | Insert | [7 empty] | 500 pip | 500 |  | 13 to 16
Stimulus type: |  | 4k CE-Chirp |  |  |  | Warning! Stimulus level exceeds recommended maximum: |  |  |  | CE-Chirp | WBand |  | 17 to 24
Patient corrected age (weeks): |  | 3 to 5 | [7 empty] | 4k CE-Chirp | 4k |  | 25 to 104
Test ear air-bone gap, dB: |  | 0 |  | 0 | [5 empty] | 2k CE-Chirp | 2k |  | >104
Non-test ear air-bone gap, dB: |  | 0 | [5 empty] | dBnHL |  | 1k CE-Chirp | 1k
dBeHL non-test BC ABR threshold: |  | 0 | [5 empty] | -3 | -61 | 500 CE-Chirp | 500
dBnHL Stimulus Level: |  | 10 |  | 15 |  | dBeHL |  | No need to mask
Biologic Nav Pro
Equation: | Stim(dBnHL) | +dBage | +RML | +ABGnt | -IA | -Nt | -Nc | Result | Offset |  |  |  | Interacoustics Eclipse
dBnoise = | 10 | 10 | 28 | 0 | 20 | 10 | -9 | 25 | +15 |  |  |  | GSI Audera
Nicolet Spirit
Variables for the current settings (dB): | [12 empty] | GN EP200/Chartr
|  | RMLlower = | 13 | 26 | [8 empty] | Neurosoft Neuro-Audio
Nt = | 10 | RMLupper = | 28 | [9 empty] | Vivosonic Integrity
Nc = | -9 | ABGnt = | 0 | [9 empty] | Custom
BC age correction = | 10 | IAs = | 0
Stimulus dBeHL = | 15 | IAn = | 54
Stimulus dBeHLnt = | -5 | IAa = | 20
|  | Max noise SPL crossing to test cochlea = | -48
Please email comments, further data, errors & omisions to: |  |  | admin@eratraining.co.uk
For ERA/ABR courses, peer reviewer training and clinical support, visit: |  |  |  | www.eratraining.co.uk
Revision history:
2011:
Transcranial transmission loss (TTL) has been renamed interaural attenuation (IA).
Values for insert & BC age-related corrections and dBnHL to dBeHL conversions reflect changes in the NHSP Early Assessment Guidelines 2011.
2012:
AC values of IA in Table 3 have been reduced by 10dB to address concerns over potential undermasking by AC. This affects when masking should be applied, the noise level and the risk of cross-masking.
BC 500Hz values in Table 4 increased by 5dB (and 1kHz values confirmed) from data on normal BC ABR thresholds in babies <6 weeks by Inga Ferm
2dB added for safety to decision on whether masking is needed.
2013:
Age ranges in Tables 4, 5 & 6 have been extended, based on available data.
BC 500Hz & 1kHz values in Table 4 increased based on data by Ferm et al 2013.
CE-Chirp stmuli have been added; eHL offsets are 5dB less than corresponding click/pip stimuli based on larger response and lower thresholds. See Table 8.
Warning against the use of low frequency BC stimuli with babies has been deleted - age-related BC offset data produced by Inga Ferm (values changed in this version)
2013b:
age range error correctred.
2018:
Noise calibration added for the Vivosonic V500, based on SPL-based noise. Age range error for max level warning corrected.
2019:
MHRA confirmed this spreadsheet does not constitute a "medical device" but users must recognise this is a guide only and take resposibility for clinical decisions.
Disclaimer: the author has produced this in good faith but can accept no responsibility for errors or inappropriate use.
2023:
NeuroAudio system added (replacing obsolete Medelec Synergy)
Changed "TDH" to "Phone"
2024:
Updated minimum IA values in table 4 (data tab) from Kariv et al 2024.
This has the effect that, below 4kHz, the calculator suggests that masking is needed at lower stimulus levels than in previous versions.
Note on A14: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking. This matters only when testing by air conduction.
Note on A15: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking.
Note on A16: Needed for deciding if masking is needed. This is the non-test ear BC threshold for the chosen stimulus type. Enter the lowest value it could be. If unknown select 0.
Note on A17: Enter the test level you are using - the calculator will then work out whether masking is needed and if so, the level of noise to use.
Note on A24: Age-related insert transducer SPL lift correction for age. Applies to stimulus & noise when applied via inserts. Data Table 5. Depends on age & stimulus.
Note on A25: Nc is the noise calibration offset for the selected system. Data Table 1. Depends on selected noise transducer
Note on A26: BC lift caused by incomplete cranial bone fusion. Data Table 6. Click data based on Webb (1993). This value is included in the equation only when BC is used.
Note on A27: For information only. Applies NHSP corrections for click BC/Insert age effects and nHL-eHL offsets.
Note on A28: Maximum stimulus level in dBeHL reaching non-test cochlea. This is helpful when deciding whether masking is necessary.
Note on C9: If "Custom", input noise calibration values in Table 1 on the Data tab.
Note on C19: Age-related correction added to stimulus level. This accounts for BC or insert age effects.
Note on C23: Level of noise that will just not totally mask the stimulus in the same ear of any subject. Used in assessing risk of cross-masking. Assumes stimulus in dBnHL and noise in dBSPL. Data Table 2.
Note on C24: Level of noise guaranteed to mask the stimulus in the same ear in all subjects. Assumes stimulus in dBnHL and noise in dBSPL. Data Table 2.
Note on C25: Air-bone gap in the non-test ear. Noise needs to be increased by this amount to overcome the effect of conductive element in the ear being masked.
Note on C26: Minimum interaural attenuation (transcanial transmission loss) for the stimulus. Data Table 3. Depends on transducer and stimulus.
Note on C27: Minimum interaural attenuation (transcanial transmission loss) for the noise. Data Table 3. Depends on transducer and stimulus.
Note on C28: Additional age-related IA (TTL). Applicable to all transducers. Data Table 4: IA depends on age & stimulus
Note on C29: Used in calculating the risk of cross-masking. Maximum noise level at the test ear cochlea in dB SPL
Note on D19: RMLupper: The level of noise guaranteed to mask the stimulus in the same ear. Assumes stimulus level in dBnHL and noise level in dBSPL
Note on E19: Air-bone gap in the non-test ear. Noise needs to be increased by this amount to overcome the effect of conductive element in the ear being masked
Note on F19: Minimum Interaural Attenuation (transcanial transmission loss) This depends on stimulus and transducer. Plus: Additional age-related IA (TTL) (incomplete skull fusion) applicable to all transducers.
Note on G19: Additional age-related correction when inserts are used as the noise transducer. Depends on age.
Note on H19: Nc is the noise calibration offset for the selected system and noise transducer. This is set by the manufacturer - there is no international standard for white noise.
Note on I19: The result has been rounded to the nearest 5dB step
Note on J19: Some systems have the option of setting the noise level as an "offset" - this is handy since changes in the stimulus level automatically change the noise level (synchronous masking)

## Notes
Notes:
1. BC & Insert age corrections taken from NHSP 2013 Early Assessment Guidelines.
2. Upper 95% (for noise level) and lower 95% (for cross masking warning) confidence level RML values taken from Lightfoot et al 2010.
3. Assumptions: ABR stimuli are calibrated to ISO 389-6 / NHSP recommended levels and noise is calibrated as manufacturers original specification.
4. Warning of stimulus level exceeding NHSP recommended maximum appears for all ages except >24 weeks.
5. Masking is needed if the stimulus level reaching the non-test cochlea exceeds the non-test BC ABR threshold (both in dBeHL).
6. Cross-masking is considered a posibility if the noise SPL reaching the test ear cochlea exceeds the corrected stimulus level (in dBnHL) reaching the test ear cochlea plus the lower EML minus 20dB (moderate risk) or 10dB (high risk)
7. CE-Chirps are assumed to have eHL corrections 5 dB lower than their click/pip counterpart.
8. The following data are needed: frequency/age corrections for tone pip insert SPL lift caused by small occluded canal volume in newborns; frequency/age IA caused by unfused cranial plates in newborns.
9. New warning of high noise levels provided if noise SPL >100dB
10. ABR systems' masking should be set to white noise, dBHL. Vivosonic: select SPL-based noise.
See also notes on Data Sheet
Don't have Microsoft Excel on your ABR system? (you must be using another computer at the moment, otherwise you wouldn't be reading this!)
You can download a free Excel viewer (you may need to ask your friendly local IT department to do this for you) for computers or as a smartphone app

## Data
Table 1: ABR system noise calibration. dB SPL noise for 0dBdial |  |  |  | Table 2: RML After Lightfoot et al (2010) |  |  | Table 3: Minimum adult IAs versus transducer & stimulus |  |  |  | Table 4: IAa. Additional age-related minimum IA versus stimulus | [5 empty] | Table 5: Insert canal volume SPL lift (used for both stimulus and noise) | [5 empty] | Table 6. BC age correction. Click data after Webb (1993). | [5 empty] | Table 7. Maximum AC stimulus levels |  |  | Table 8. nHL-eHL offsets
System | Nc Phones | Nc Ins/2cc | Nc Ins/711 | Stimulus | RMLupper | RMLlower | Stimulus | IA phones | IA inserts | IA BC | Corrected age (weeks) | Wband | 4k | 2k | 1k | 500 | Corrected age (weeks) | Wband | 4k | 2k | 1k | 500 | Corrected age (weeks) | Wband | 4k | 2k | 1k | 500 | Stimulus | Phones | Inserts | Stimulus | dB
Biologic Nav Pro | 21 | 19 | 28 | Click | 33 | 18 | Click | 38 | 45 | 0 | <-3 | 20 | 20 | 10 | 0 | 0 | <-3 | 10 | 10 | 5 | 5 | 5 | <-3 | 12 | 10 | 5 | 20 | 25 | Click | 100 | 85 | Click | 5
Interacoustics Eclipse | 0 | -9 | 0 | 4k pip | 28 | 13 | 4k pip | 42 | 54 | 0 | -3 to -1 | 20 | 20 | 10 | 0 | 0 | -3 to -1 | 10 | 10 | 5 | 5 | 5 | -3 to -1 | 10 | 10 | 5 | 20 | 25 | 4k pip | 105 | 85 | 4k pip | 10
GSI Audera | 0 | -9 | 0 | 2k pip | 33 | 13 | 2k pip | 35 | 44 | 0 | 0 to 2 | 20 | 20 | 10 | 0 | 0 | 0 to 2 | 10 | 10 | 5 | 5 | 5 | 0 to 2 | 8 | 10 | 5 | 20 | 25 | 2k pip | 110 | 95 | 2k pip | 10
Nicolet Spirit | 20 | 17 | 26 | 1k pip | 28 | 13 | 1k pip | 37 | 46 | 0 | 3 to 5 | 20 | 20 | 10 | 0 | 0 | 3 to 5 | 10 | 10 | 5 | 5 | 5 | 3 to 5 | 6 | 10 | 5 | 20 | 25 | 1k pip | 115 | 100 | 1k pip | 15
GN EP200/Chartr | 0 | 0 | 9 | 500 pip | 33 | 18 | 500 pip | 35 | 40 | 0 | 6 to 8 | 20 | 20 | 10 | 0 | 0 | 6 to 8 | 10 | 10 | 5 | 5 | 5 | 6 to 8 | 5 | 10 | 5 | 20 | 25 | 500 pip | 110 | 100 | 500 pip | 20
Neurosoft Neuro-Audio | 0 | -9 | 0 | CE-Chirp | 33 | 18 | CE-Chirp | 38 | 45 | 0 | 9 to 12 | 20 | 20 | 10 | 0 | 0 | 9 to 12 | 10 | 10 | 5 | 5 | 5 | 9 to 12 | 3 | 10 | 5 | 20 | 25 | CE-Chirp | 100 | 85 | CE-Chirp | 0
Vivosonic Integrity | 0 | -9 | 0 | 4k CE-Chirp | 28 | 13 | 4k CE-Chirp | 42 | 54 | 0 | 13 to 16 | 15 | 15 | 0 | 0 | 0 | 13 to 16 | 5 | 5 | 0 | 0 | 0 | 13 to 16 | 0 | 5 | 0 | 15 | 20 | 4k CE-Chirp | 105 | 85 | 4k CE-Chirp | 5
Custom | 0 | -9 | 0 | 2k CE-Chirp | 33 | 13 | 2k CE-Chirp | 35 | 44 | 0 | 17 to 24 | 5 | 15 | 0 | 0 | 0 | 17 to 24 | 5 | 5 | 0 | 0 | 0 | 17 to 24 | 0 | 5 | 0 | 15 | 20 | 2k CE-Chirp | 110 | 95 | 2k CE-Chirp | 5
Notes: Based on a sample of available systems (6cc NBS 9A and 2cc HA-2 couplers) or manufacturer information and assumes the noise calibration of your system is similar - check with your calibration agent. Ensure any "offset" in the test protocol is zero when calibrating the noise level and that the transducer type being used matches the type for the calibration file. Custom values may be entered if desired. |  |  |  | 1k CE-Chirp | 28 | 13 | 1k CE-Chirp | 37 | 46 | 0 | 25 to 104 | 0 | 0 | 0 | 0 | 0 | 25 to 104 | 0 | 0 | 0 | 0 | 0 | 25 to 104 | 0 | 0 | 0 | 10 | 15 | 1k CE-Chirp | 115 | 100 | 1k CE-Chirp | 10
|  |  |  | 500 CE-Chirp | 33 | 18 | 500 CE-Chirp | 35 | 40 | 0 | >104 | 0 | 0 | 0 | 0 | 0 | >104 | 0 | 0 | 0 | 0 | 0 | >104 | 0 | 0 | 0 | 0 | 0 | 500 CE-Chirp | 110 | 100 | 500 CE-Chirp | 15
|  |  |  | Notes: Relative masking level. Based on subjective masking data in adults. Chirps: assume same as click/pip data. |  |  | Notes: Data on how IA changes with frequency is incomplete. Further research is needed. Source: Phones: average of a number of published studies for pure tones, lower 95% confidence level. Click value taken as mean of 2k & 4kHz. Inserts: Munro & Contractor, 2010, minimum of range in 15 subjects. BC: 0dB taken as conservative value. 2012: Phones & insert IA values have been reduced by 10dB at all frequencies to address concerns over undermasking by AC. |  |  |  | Notes: IA is assumed to be greater in immature skulls because of incomplete fusion of cranial bones. Further research is needed to define how these values change with age and frequency. Applicable to clicks, pips & chirps. Wband = clicks & CE chirps. Tonal data up to 1 year from Kariv et al 2024. | [5 empty] | Notes. From NHSP Early Assessment Guidelines, 2013. Further research is needed to further refine these data. Applicable to clicks, pips & chirps. Wband = clicks & CE chirps. | [5 empty] | Notes. Click data from Webb (1993), extrapolated using exponential function. 1k & 500Hz: Data from study by Ferm 2013. Applicable to clicks, pips & chirps. Wband = clicks & CE chirps. | [5 empty] | Notes: From NHSP Early Assessment Guidelines, 2013. The spreadsheet applies these levels up to 24 weeks corrected age. |  |  | Notes: Chirp values are 5dB less than click/pip values. Data derived from pip/chirp study by Ferm, Lightfoot & Stevens 2013.
Note on B10: Enter your system's noise dBSPL above the dial level, measured in an NBS 9A 6cc coupler
Note on C10: Enter your system's noise dBSPL above the dial level, measured in an HA-2 2cc coupler
Note on D10: Enter your system's noise dBSPL above the dial level, measured in a 711 occluded ear simulator then subtract 9dB and enter that value in the cell to the left