## Set-up
Spreadsheet to calculate audiometric target SLM values
Assumptions:
a) That ISO 389 series reference values are used (NHSP recommended values used for unavailable figures)
b) That a fixed stimulus Hearing (dial) Level is used at all frequencies for a given transducer / stimulus type
c) That the same microphone is used for the IEC 60318-3 & -5 (9A and 2cc couplers)
d) That for non-NPL calibrated systems, the frequency response of the measuring system is flat except for the mic/mastoid
Note: Spreadsheet contains typical values - replace numbers in mustard-coloured cells with values specific to your system
Hovering the mouse over a cell with a red corner reveals a help message or explanation - try this example --------------------> | [10 empty] | hover here
Step 1: Specify the Hearing Levels at which each transducer type is to be calibrated for pure tones.
Measuring device |  | IEC 60318-1 |  | IEC 60318-3 |  | IEC 60318-4 |  | IEC 60318-5 |  | IEC 60318-6
dBHL |  | 60 |  | 60 |  | 60 |  | 60 |  | 30
Step 2: Specify the Hearing Levels at which each transducer type is to be calibrated for clicks & tone pips.
Measuring device |  | IEC 60318-1 |  | IEC 60318-3 |  | IEC 60318-4 |  | IEC 60318-5 |  | IEC 60318-6
dBHL |  | 80 |  | 80 |  | 80 |  | 80 |  | 40
Step 3: Enter the values representing the sensitivity of the IEC 60318-1 microphone at each frequency
Freq (Hz) | 125 | 250 | 500 | 750 | 1000 | 1500 | 2000 | 3000 | 4000 | 6000 | 8000
dB | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.3 | 0.6
Step 4: Enter the values representing the sensitivity of the IEC 60318-3 & -5 microphone at each frequency
Freq (Hz) | 125 | 250 | 500 | 750 | 1000 | 1500 | 2000 | 3000 | 4000 | 6000 | 8000
dB | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.5 | 0.5 | 0.5 | -1.5
Step 5: Enter the values representing the sensitivity of the IEC 60318-4 microphone at each frequency
Freq (Hz) | 125 | 250 | 500 | 750 | 1000 | 1500 | 2000 | 3000 | 4000 | 6000 | 8000
dB | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.5 | 1
Step 6: Indicate whether you have NPL targets or manufacturer sensitivity values for your artificial mastoid
Enter 1 for NPL, 2 for mastoid sensitivity values----> |  |  |  |  | 1 | You've selected NPL mastoid target values - go to step 9
Step 7: Enter the artificial mastoid overall sensitivity at 1 kHz in dB Re: µV/µN or V/N here:----> | [9 empty] | -17.2 |  | Force Sensitivity quoted in mV/N?
[12 empty] | Enter value here -> | 138
Step 8: Enter the artificial mastoid frequency sensitivity relitive to 1 kHz (then skip step 9 & go to step 10) | [11 empty] | dB Re: V/N is: | -17.2024
Freq (Hz) |  | 250 | 500 | 750 | 1000 | 1500 | 2000 | 3000 | 4000 | 6000 | 8000
dB |  | -1 | -0.7 | -0.5 | 0 | 1 | 2.5 | 1 | -6 | -9.5 | -8
Step 9: Enter the NPL bone target values
Freq (Hz) |  | 250 | 500 | 750 | 1000 | 1500 | 2000 | 3000 | 4000 | 6000 | 8000
dB |  | 78 | 69 | 60 | 55 | 50 | 45 | 44 | 45 | 43 | 43
Step 10: SLM target values are now shown on PTA & ABR worksheets respectively
Send queries, comments or corrections to Guy Lightfoot (guy.lightfoot@nhs.net)
Disclaimer: The author can accept no responsibility for errors or omissions to this spreadsheet
Note on A19: If your system is calibrated by NPL, enter the system correction values for this microphone
Note on A23: If your system is calibrated by NPL, enter the system correction values for this microphone
Note on A27: If your system is calibrated by NPL, enter the system correction values for this microphone
Note on A34: Ignore this step if you have selected "1" in step 6
Note on A36: Ignore this step if you have selected "1" in step 6
Note on A40: Ignore this step if you selected "2" in step 6
Note on C12: Wide-band artificial ear e.g. B&K 4153
Note on C16: Wide-band artificial ear e.g. B&K 4153
Note on E12: 6cc / NBS-9A coupler e.g. B&K 4152
Note on E16: 6cc / NBS-9A coupler e.g. B&K 4152
Note on F38: By definition, this value should be zero
Note on G12: IEC 60711 occluded ear simulator e.g. B&K 4157
Note on G16: IEC 60711 occluded ear simulator e.g. B&K 4157
Note on I12: 2cc Coupler / HA-2 e.g. B&K DB0138
Note on I16: 2cc Coupler / HA-2 e.g. B&K DB0138
Note on K12: IEC 60373 artificial mastoid e.g. B&K 4930
Note on K13: If your system is calibrated by NPL, this must be the level used in that calibration - it will be 30
Note on K16: IEC 60373 artificial mastoid e.g. B&K 4930
Note on K34: If you are using a B&K 4930, the figure is likely to be typically -18
Note on L9: Example of a help message

## PTA
Pure Tone Air Conduction |  |  |  |  | Pure Tone Bone Conduction
IEC 60318-1
Freq | Mic Sens | HL | RET | Target | IEC 60373 | For use with manufacturers mastoid sensitivity values
125 | 0 | 60 | 45 | 105 | Freq | RETFL | Sens 1k | Sens Re 1k | HL | Target | Target
250 | 0 | 60 | 27 | 87 | 250 | 67 | -17.2 | -1 | 30 | 78.8 | 79
500 | 0 | 60 | 13.5 | 73.5 | 500 | 58 | -17.2 | -0.7 | 30 | 70.1 | 70
750 | 0 | 60 | 9 | 69 | 750 | 48.5 | -17.2 | -0.5 | 30 | 60.8 | 61
1000 | 0 | 60 | 7.5 | 67.5 | 1000 | 42.5 | -17.2 | 0 | 30 | 55.3 | 55
1500 | 0 | 60 | 7.5 | 67.5 | 1500 | 36.5 | -17.2 | 1 | 30 | 50.3 | 50
2000 | 0 | 60 | 9 | 69 | 2000 | 31 | -17.2 | 2.5 | 30 | 46.3 | 46
3000 | 0 | 60 | 11.5 | 71.5 | 3000 | 30 | -17.2 | 1 | 30 | 43.8 | 44
4000 | 0 | 60 | 12 | 72 | 4000 | 35.5 | -17.2 | -6 | 30 | 42.3 | 42
6000 | 0.3 | 60 | 16 | 76.3 | 6000 | 40 | -17.2 | -9.5 | 30 | 43.3 | 43
8000 | 0.6 | 60 | 15.5 | 76.1 | 8000 | 40 | -17.2 | -8 | 30 | 44.8 | 45
TDH / IEC 60318-3
Freq | Mic Sens | HL | RET | Target | IEC 60373 | For use with NPL certificate
125 | 0 | 60 | 45 | 105 | Freq | Target
250 | 0 | 60 | 25.5 | 85.5 | 250 | 78
500 | 0 | 60 | 11.5 | 71.5 | 500 | 69
750 | 0 | 60 | 7.5 | 67.5 | 750 | 60
1000 | 0 | 60 | 7 | 67 | 1000 | 55
1500 | 0 | 60 | 6.5 | 66.5 | 1500 | 50
2000 | 0 | 60 | 9 | 69 | 2000 | 45
3000 | 0.5 | 60 | 10 | 70.5 | 3000 | 44
4000 | 0.5 | 60 | 9.5 | 70 | 4000 | 45
6000 | 0.5 | 60 | 15.5 | 76 | 6000 | 43
8000 | -1.5 | 60 | 13 | 71.5 | 8000 | 43
ER-3A Insert / IEC 60318-4
Freq | Mic Sens | HL | RET | Target
125 | 0 | 60 | 28 | 88
250 | 0 | 60 | 17.5 | 77.5
500 | 0 | 60 | 9.5 | 69.5
750 | 0 | 60 | 6 | 66
1000 | 0 | 60 | 5.5 | 65.5
1500 | 0 | 60 | 9.5 | 69.5
2000 | 0 | 60 | 11.5 | 71.5
3000 | 0 | 60 | 13 | 73
4000 | 0 | 60 | 15 | 75
6000 | 0.5 | 60 | 16 | 76.5
8000 | 1 | 60 | 15.5 | 76.5
ER-3A Insert / IEC 60318-5
Freq | Mic Sens | HL | RET | Target
125 | 0 | 60 | 26 | 86
250 | 0 | 60 | 14 | 74
500 | 0 | 60 | 5.5 | 65.5
750 | 0 | 60 | 2 | 62
1000 | 0 | 60 | 0 | 60
1500 | 0 | 60 | 2 | 62
2000 | 0 | 60 | 3 | 63
3000 | 0.5 | 60 | 3.5 | 64
4000 | 0.5 | 60 | 5.5 | 66
6000 | 0.5 | 60 | 2 | 62.5
8000 | -1.5 | 60 | 0 | 58.5
1
Note on A3: Wide-band artificial ear e.g. B&K 4153
Note on A17: 6cc / NBS-9A coupler e.g. B&K 4152
Note on A31: IEC 60711 occluded ear simulator e.g. B&K 4157
Note on A45: 2cc Coupler / HA-2 e.g. B&K DB0138
Note on D4: From ISO 389 series
Note on D18: From ISO 389 series
Note on D32: From ISO 389 series
Note on D46: From ISO 389 series
Note on G4: Will become IEC 60318-6 Artificial Mastoid e.g. B&K 4930
Note on G14: Technical limitations with the design of bone vibrators at 6 kHz and above make clinical tests at these frequencies inadvisable.
Note on G15: Technical limitations with the design of bone vibrators at 6 kHz and above make clinical tests at these frequencies inadvisable.
Note on G18: Will become IEC 60318-6 Artificial Mastoid e.g. B&K 4930
Note on G28: Technical limitations with the design of bone vibrators at 6 kHz and above make clinical tests at these frequencies inadvisable.
Note on G29: Technical limitations with the design of bone vibrators at 6 kHz and above make clinical tests at these frequencies inadvisable.
Note on H5: From ISO 389-3
Note on I5: Mastoid voltage sensitivity (dB Re 1V/N at 1 kHz)
Note on J5: voltage sensitivity of mastoid in dB Re: value at 1 kHz
Note on K5: Dail level in dBHL at which calibration performed
Note on L5: Target SLM reading (in dB Re: 1uV)
Note on M5: Target rounded to nearest dB

## ABR Ck-Pip
ABR Air Conduction |  |  |  |  | ABR Bone Conduction
IEC 60318-1 |  |  |  |  | IEC 60373 | For use with manufacturers mastoid sensitivity values | [5 empty] | Targets for 80dBVFL |  | For use with mastoid sensitivity values
Freq | Mic Sens | HL | RET | Target | Freq | RETFL | 1 k Sens | Sens Re 1k | HL | Target | Target | Freq | dB VFL | 1 k Sens | Sens Re 1k | Target | Target
500 | 0 | 80 | 23 | 103 | 500 | 69.5 | -17.2 | -0.7 | 40 | 91.6 | 92 | 500 | 80 | -17.2 | -0.7 | 62.1 | 62
1000 | 0 | 80 | 18.5 | 98.5 | 1000 | 58.5 | -17.2 | 0 | 40 | 81.3 | 81 | 1000 | 80 | -17.2 | 0 | 62.8 | 63
2000 | 0 | 80 | 25 | 105 | 2000 | 47.5 | -17.2 | 2.5 | 40 | 72.8 | 73 | 2000 | 80 | -17.2 | 2.5 | 65.3 | 65
4000 | 0 | 80 | 27.5 | 107.5 | 4000 | 53 | -17.2 | -6 | 40 | 69.8 | 70 | 4000 | 80 | -17.2 | -6 | 56.8 | 57
Click | 0 | 80 | 31 | 111 | Click | 51.5 | -17.2 | 0 | 40 | 74.3 | 74 | Click | 80 | -17.2 | 0 | 62.8 | 63
TDH / IEC 60318-3 |  |  |  |  | IEC 60373 | For use with NPL certificate | [5 empty] | Targets for 80dBVFL |  | For use with NPL certificate
Freq | Mic Sens | HL | RET | Target | Freq | RETFL | Cal level | NPL data | HL | Target | Target | Freq | dB VFL | Cal level | NPL data | Target | Target
500 | 0 | 80 | 23 | 103 | 500 | 69.5 | 88 | 69 | 40 | 90.5 | 91 | 500 | 80 | 88 | 69 | 61 | 61
1000 | 0 | 80 | 18.5 | 98.5 | 1000 | 58.5 | 72.5 | 55 | 40 | 81 | 81 | 1000 | 80 | 72.5 | 55 | 62.5 | 63
2000 | 0 | 80 | 25 | 105 | 2000 | 47.5 | 61 | 45 | 40 | 71.5 | 72 | 2000 | 80 | 61 | 45 | 64 | 64
4000 | 0.5 | 80 | 27.5 | 108 | 4000 | 53 | 65.5 | 45 | 40 | 72.5 | 73 | 4000 | 80 | 65.5 | 45 | 59.5 | 60
Click | 0 | 80 | 31 | 111 | Click | 51.5 | 72.5 | 55 | 40 | 74 | 74 | Click | 80 | 72.5 | 55 | 62.5 | 63
ER-3A Insert / IEC 60318-4
Freq | Mic Sens | HL | RET | Target
500 | 0 | 80 | 23.5 | 103.5
1000 | 0 | 80 | 21.5 | 101.5
2000 | 0 | 80 | 28.5 | 108.5
4000 | 0 | 80 | 32.5 | 112.5
Click | 0 | 80 | 35.5 | 115.5
ER-3A Insert / IEC 60318-5
Freq | Mic Sens | HL | RET | Target
500 | 0 | 80 | 19.5 | 99.5
1000 | 0 | 80 | 16 | 96
2000 | 0 | 80 | 20 | 100
4000 | 0.5 | 80 | 23 | 103.5
Click | 0 | 80 | 26.5 | 106.5
1
Note on A3: Wide-band artificial ear e.g. B&K 4153
Note on A11: 6cc / NBS-9A coupler e.g. B&K 4152
Note on A19: IEC 60711 occluded ear simulator e.g. B&K 4157
Note on A27: 2cc Coupler / HA-2 e.g. B&K DB0138
Note on D4: From ISO 389 series
Note on D12: From ISO 389 series
Note on D20: From ISO 389 series
Note on D28: From ISO 389 series
Note on G3: Will become IEC 60318-6 Artificial Mastoid e.g. B&K 4930
Note on G11: Will become IEC 60318-6 Artificial Mastoid e.g. B&K 4930
Note on H4: RETFL in dB Re: 1uN from ISO 389-6 (click) and NHSP data (pips)
Note on H12: RETFL in dB Re: 1uN from ISO 389-6 (click) and NHSP data (pips)
Note on I4: Mastoid voltage sensitivity (dB Re 1V/N at 1 kHz)
Note on I12: The force level (RETFL+HL) used during the NPL calibration
Note on J4: voltage sensitivity of mastoid in dB Re: value at 1 kHz
Note on J12: Target values from the pure tone NPL bone conduction certificate
Note on K4: Dail level in dBHL at which calibration performed
Note on K12: Dail level in dBHL at which calibration performed
Note on L4: Target SLM reading (in dB Re: 1uV)
Note on L12: Target SLM reading (in dB)
Note on M4: Target rounded to nearest dB
Note on M12: Target rounded to nearest dB
Note on Q4: Mastoid voltage sensitivity (dB Re 1V/N at 1 kHz)
Note on Q12: The force level (RETFL+HL) used during the NPL calibration
Note on R4: voltage sensitivity of mastoid in dB Re: value at 1 kHz
Note on R12: Target values from the pure tone NPL bone conduction certificate
Note on S4: Target SLM reading (in dB Re: 1uV)
Note on S12: Target SLM reading (in dB)
Note on T4: Target rounded to nearest dB
Note on T12: Target rounded to nearest dB

## ABR Chirps
ABR Air Conduction |  |  |  |  | ABR Bone Conduction
IEC 60318-1 |  |  |  |  | IEC 60373 | For use with manufacturers mastoid sensitivity values | [5 empty] | Targets for 80dBVFL |  | For use with mastoid sensitivity values
Freq | Mic Sens | HL | RET | Target | Freq | RETFL | 1 k Sens | Sens Re 1k | HL | Target | Target | Freq | dB VFL | 1 k Sens | Sens Re 1k | Target | Target
500 | 0 | 80 | 26 | 106 | 500 | 75 | -17.2 | -0.7 | 40 | 97.1 | 97 | 500 | 80 | -17.2 | -0.7 | 62.1 | 62
1000 | 0 | 80 | 21.5 | 101.5 | 1000 | 61.5 | -17.2 | 0 | 40 | 84.3 | 84 | 1000 | 80 | -17.2 | 0 | 62.8 | 63
2000 | 0 | 80 | 27 | 107 | 2000 | 50 | -17.2 | 2.5 | 40 | 75.3 | 75 | 2000 | 80 | -17.2 | 2.5 | 65.3 | 65
4000 | 0 | 80 | 30 | 110 | 4000 | 55.5 | -17.2 | -6 | 40 | 72.3 | 72 | 4000 | 80 | -17.2 | -6 | 56.8 | 57
Click | 0 | 80 | 27.5 | 107.5 | Click | 51.5 | -17.2 | 0 | 40 | 74.3 | 74 | Click | 80 | -17.2 | 0 | 62.8 | 63
TDH / IEC 60318-3 |  |  |  |  | IEC 60373 | For use with NPL certificate | [5 empty] | Targets for 80dBVFL |  | For use with NPL certificate
Freq | Mic Sens | HL | RET | Target | Freq | RETFL | Cal level | NPL data | HL | Target | Target | Freq | dB VFL | Cal level | NPL data | Target | Target
500 | 0 | 80 | 26 | 106 | 500 | 75 | 88 | 69 | 40 | 96 | 96 | 500 | 80 | 88 | 69 | 61 | 61
1000 | 0 | 80 | 21.5 | 101.5 | 1000 | 61.5 | 72.5 | 55 | 40 | 84 | 84 | 1000 | 80 | 72.5 | 55 | 62.5 | 63
2000 | 0 | 80 | 27 | 107 | 2000 | 50 | 61 | 45 | 40 | 74 | 74 | 2000 | 80 | 61 | 45 | 64 | 64
4000 | 0.5 | 80 | 30 | 110.5 | 4000 | 55.5 | 65.5 | 45 | 40 | 75 | 75 | 4000 | 80 | 65.5 | 45 | 59.5 | 60
Click | 0 | 80 | 27.5 | 107.5 | Click | 51.5 | 72.5 | 55 | 40 | 74 | 74 | Click | 80 | 72.5 | 55 | 62.5 | 63
ER-3A Insert / IEC 60318-4
Freq | Mic Sens | HL | RET | Target
500 | 0 | 80 | 26.5 | 106.5 |  | Note:
1000 | 0 | 80 | 24.5 | 104.5 |  | The chirps referred to here are those specified by Elberling, C., and Don, M. (2010).
2000 | 0 | 80 | 30.5 | 110.5 |  | “A direct approach for the design of chirp stimuli used for the recording of auditory brainstem responses,”
4000 | 0 | 80 | 35 | 115 |  | See Figs 1 & 2. J. Acoust. Soc. Am. 128, 2955-2964
Click | 0 | 80 | 35 | 115
ER-3A Insert / IEC 60318-5
Freq | Mic Sens | HL | RET | Target
500 | 0 | 80 | 22.5 | 102.5
1000 | 0 | 80 | 18.5 | 98.5
2000 | 0 | 80 | 22 | 102
4000 | 0.5 | 80 | 25 | 105.5
Click | 0 | 80 | 22 | 102
1
Note on A3: Wide-band artificial ear e.g. B&K 4153
Note on A11: 6cc / NBS-9A coupler e.g. B&K 4152
Note on A19: IEC 60711 occluded ear simulator e.g. B&K 4157
Note on A27: 2cc Coupler / HA-2 e.g. B&K DB0138
Note on D4: See note on this page
Note on D12: See note on this page
Note on D20: See note on this page
Note on D28: See note on this page
Note on G3: Will become IEC 60318-6 Artificial Mastoid e.g. B&K 4930
Note on G11: Will become IEC 60318-6 Artificial Mastoid e.g. B&K 4930
Note on H4: RETFL in dB Re: 1uN from ISO 389-6 (click) and NHSP data (pips)
Note on H12: RETFL in dB Re: 1uN from ISO 389-6 (click) and NHSP data (pips)
Note on I4: Mastoid voltage sensitivity (dB Re 1V/N at 1 kHz)
Note on I12: The force level (RETFL+HL) used during the NPL calibration
Note on J4: voltage sensitivity of mastoid in dB Re: value at 1 kHz
Note on J12: Target values from the pure tone NPL bone conduction certificate
Note on K4: Dail level in dBHL at which calibration performed
Note on K12: Dail level in dBHL at which calibration performed
Note on L4: Target SLM reading (in dB Re: 1uV)
Note on L12: Target SLM reading (in dB)
Note on M4: Target rounded to nearest dB
Note on M12: Target rounded to nearest dB
Note on Q4: Mastoid voltage sensitivity (dB Re 1V/N at 1 kHz)
Note on Q12: The force level (RETFL+HL) used during the NPL calibration
Note on R4: voltage sensitivity of mastoid in dB Re: value at 1 kHz
Note on R12: Target values from the pure tone NPL bone conduction certificate
Note on S4: Target SLM reading (in dB Re: 1uV)
Note on S12: Target SLM reading (in dB)
Note on T4: Target rounded to nearest dB
Note on T12: Target rounded to nearest dB